## PTS
Technical Specifications of PTS [Pneumatic Tube System]
Please quote all the models including the state of art equipment which meets the basic minimum specifications.
Model Name:
Make:
Year of Introduction:
Probable end of support:
SN. | Description of Item | Compliance Yes/No | Remarks
1.1 | Pneumatic tube transport system of 160mm pipe (outer diameter) for high speed (minimum transfer rate of 2 m/s to maximum transfer rate of 6 m/s) spontaneous automated vertical and horizontal transports of blood samples, pharmaceuticals, documents etc.
1.2 | Blood samples should be transported at slow speed of 2 to 3 m/s with help of frequency controller (VFD)/ Station software . Empty carrier returns and the transport of other goods should be with increased speed of up to 5 to 6 m/s.
1.3 | The system should be capable to transporting various materials of minimum 4 kg
1.4 | The equipment must be supplied with relevant Declarations of Conformity to certify compliance with the EMC directive 89/336/EEC-92/31/EEC and the Machinery Safety Directive 89/392/EEC91/368/EEC-93/44/EEC.DIN regulations: Tubing (PVC-U) must meet DIN 6660/6661 regulations - DIN8061/62 Group B1 specification. Fire: The equipment must meet the inflammable category B1, difficult inflammable. Fire: self-extinguishing: Where tubing passes through a wall, floor, ceiling or other barrier, the fire rating of the barrier is not accepted to be reduced. This must be achieved by equipping the system on these places, with crushing/self-extinguishing fire sleeves/collars that are certified by a notified body and are compliant with the local regulations. The Bidder must submit the factory Test report, dispatch Lot details and the proof of country of origin.
1.5 | All items and accessories should be of Imported (Except PC and nut and bolts) and comply with the below specifications
2 | Main Controller: The system must be fully automatic computer-controlled. The PC serves as a control and supervising centre that controls and monitors all transmissions within the whole system on a continuous basis.
2.1 | The sending process has to be indicated on display devices. The device also has to provide information to find the cause of a system malfunction. Customer‐specific data such as the system's layout, target numbers, target names, arrival signals, and priority and special functions must be selectable on site without change or external reprogramming of memory devices. An integrated uninterrupted power supply must provide the requested system back up time to paste all date before shutting down the system, so after power source is providing energy again an automatic system start will occur and the system status will get back in the operation mode as before. Chosen targets at stored containers must be kept in the memory to proceed them automatically after power failure restart.
2.2 | All components of the pneumatic tube conveyor should be constantly monitored; the operating software has to be based on action reaction control for any device. The status of each device should be checked by the master control unit. A test program must be included to automatically check, move and supervise all of the system's devices, or specific selected devices, by access via service code from the Station control panel.
2.3 | During both normal operation and testing, all devices should inform the master control unit that the selected functional position has been reached. The system should be designed in such a way that it does not allow the unobserved pivoting of devices.
2.4 | Automatic purge cycle after transmission time surpassed, quantity limitation per day should be included.
2.5 | Main Controller should include the following:
2.5a | Main Control Unit Hardware (Viz: CPU, Monitor Keyboard, Mouse, and UPS)
2.5b | Software package for Central Controller includes Supervision; Real time monitoring; Priority settings; Network Graphic visualization; System access and controls all devices on the network; Traffic flow analysis; Logging of each transaction and error; Carrier Tracking Password Access controls; Pro-remote accessibility; Divert Carrier function in case of absence; Various Priority Settings for urgent and emergency transports; Print module for reports and statistics, analysis reports; User profiles and rights; The Control PC should be PC with CPU with INTEL I7 with SMPS Cabinet 3.4GHz/8GB Ram; 1 TB HDD; RS232 Port (atleast 2) USB Ports (atleast 2); PCI Slot for ISDN/Modem Card Data Port; Ethernet 100/1000 Lan Card Standard Keyboard Standard Optical mouse; 22” flat LCD Monitor; Standard Mono laser A4 printer; Original licensed MS-Office2016, and Windows/Linux and OS Antivirus licensed copy with and necessary softwares, UPS etc.
2.5c | Power Supply Unit with components
2.5d | Additional Power Supply Unit with components
2.5e | Power Supply connecting cable from Isolator panel and MCCB/RCCB for Three Phase power supply
2.5f | The control software continuously monitors all sensors, switches, motors and other components, and gives early warning should the performance of any component start to degrade. This enables maintenance to be carried out prior to absolute failure and keeps system downtime to an absolute minimum
2.5g | The controller Should be Integrated with the Building Management System to warn of an alarm status.
2.5h | The controller will be connected to the fire alarm system to enable the pneumatic tube system to be automatically shut down in the event of a fire. It can be selected that the current dispatch within a system will continue to its destination before shut down occurs. No new dispatch will be accepted after alarm has been triggered until alarm status has returned to normal
3 | Side Channel Blower: 3 to 5.5 Kw unidirectional side channel blower confirms to EC machines directives 98/37/EC and low voltage director 73/23/EEC f.; DIN EN292 for safety of machines; EN60034-1 / DIN VDE 0530 Part 1 for rotating electrical machines; EN 60034-5 / DIN VDE 0530-5 classification of degree of protection DIN EN 60204 for safety of machine electrical equipment of machine Part 1 of Electric Motors and DIN VDE 0-1 insulation coordination for equipment within low voltage systems. Complete with silencer, filters, dampers & installation accessories and air switch device.Automatic activation through centralized control system. The Bidder will supply and install the VFD for variable speed control if the feature is not available through software .
3.1 | Each blower should be provided with Frequency Converter or VFD for Control of slow speed for sensitive laboratory samples by frequency control of Compressor. The blower should be set up to 75Hz with the help of Frequency Converter.
3.2 | It should be provided with all the mounting accessories
4 | Top Load/ Front Load and Pass through Stations: : The receiving and sending stations for automatic receiving and sending of materials in the carriers. To conform to EU Director for Electromagnetic tolerance 2004/108/EC; EEC directive for low voltage devices 93/68/EEC; EEC machines directors 98/37/EEC. To have ergonomically adapted sending and receiving heights. Wall mountable stations; Each device equipped with state-of-art maintenance free gear drive mechanism & self-adjusting optical seals for moving parts for noise-less operations. Each station should be provided with carrier receiving basket with cushion, carrier rack. Each stations should be equipped with RFID reading device.
4.1 | The Pneumatic Station should be able to send and receive containers.
4.2 | Inserting a container into the Pneumatic Station and selecting a target number should be possible. The container should be loaded on the top/ front side of the Pneumatic Station.
4.3 | The Pneumatic Station should be made for longer life, maintenance free mechanism, with self-adjusting optical switches, with self-adjusting maintenance free gaskets for noise less operations, contact less sensoring of the unit positions. At terminal / end stations where air-exits should be installed 1D special filter housing with H13/H14 plus prefilter; and at blowers should be installed 2D special filter housing with T-connect and H13/14 plus prefilter.
4.4 | The Pneumatic Station should have soft landing facility for arriving container to protect samples. Provided with container rack and receiving basket with cushion from the same make and country of origin .
5 | Compact End / Front loading Station: The receiving and sending stations for automatic receiving and sending of materials in the carriers. To conform to EU Director for Electromagnetic tolerance 2004/108/EC; EEC directive for low voltage devices 93/68/EEC; EEC machines directors 98/37/EEC. To have ergonomically adapted sending and receiving heights. Wall mountable stations; Each device equipped with state-of-art maintenance free gear drive mechanism & self adjusting optical seals for moving parts for noise-less operations. Each station should be provided with carrier receiving basket with cushion, carrier rack. Each stations should be equipped with RFID reading device.The Pneumatic Station should be able to send and receive containers from the same unit.
5.3 | It should be Microprocessor‐controlled.
5.5 | It should be designed as a fully automatic dispatch and receiving unit and can only be used as end station.
5.7 | It should be built in a way that after a power failure it is self‐examining and if necessary self‐repairing.
5.8 | The Pneumatic Station should be for longer life, maintenance free mechanism, with self-adjusting optical switches, with self-adjusting maintenance free gaskets for noise less operations, contact less sensoring of the unit positions. At terminal / end stations where air-exits should be installed 1D special filter housing with H13/H14 plus prefilter; and at blowers should be installed 2D special filter housing with T-connect and H13/14 plus prefilter.
5.9 | It should have cushioned soft landing facility for arriving containers to protect samples.
5.1 | It should be provided with container rack & PVC Slide bend, sliced from the top for soft landing of the samples from same make and country of origin.
6 | Horizontal Lab Receiving Station:
6.1 | The station Device for installation in laboratories with much reduced space capacity and always taking in consideration the ergonomically aspect. Conforms to EU Director for Electromagnetic tolerance 2004/108/EC; EEC directive for low voltage devices 93/68/EEC; EEC machines directors 98/37/EEC. With built-in special turn- around function, the carriers slow down, turn from a vertical position in the station to a horizontal position and roll horizontally in the arrival area. Receiving area at 93degrees with 2000mmL; the device is compact size and connected to dedicated and independent powerlines. The device is wall / table top mountable; Each device should be equipped with state-of-art maintenance free gear drive mechanism & self-adjusting optical seals for moving parts for noise-less operations.
6.2 | The Pneumatic Station should be made for longer life, maintenance free mechanism, with self-adjusting optical switches, with self-adjusting maintenance free gaskets for noise less operations, contact less sensoring of the unit positions. At terminal / end stations where air-exits should be installed 1D special filter housing with H13/H14 plus prefilter; and at blowers should be installed 2D special filter housing with T-connect and H13/14 plus prefilter.
7 | Optical tube switch Modular switch provided with built-in various devices, portable in nature, easily replaceable, com-connectable to various devices through the system with single control cable of high grade composite with grounding, power and data-com all three in one and fully screen and use grounded system concept to minimize electromagnetic and radio-frequency interference. Instrumentation wire type 2 x 2 x 0,22 x 1,9², Compatible with system on low voltage supply of 36V for human protection.
8 | Tubes & Bends 160mm: Grey UPVC or PVC-U tubes and bends with Wall thickness 3.2mm, Radius 800mm. Special Imported U-PVC forwarding bends conforming to DIN / CE standards for characteristics; with following features - Physical tensile strength: 55N/mm3; E- module:3000N/mm3; Resistance to impact: High resistance at 20 degrees Celsius without breaking; Thermal Coefficient of Linea Expansion: 80 x 10-6 K-1; Heat conductivity:0.16 W/mk; Inherent stability up to: 60 degrees Celsius; Electric surface resistance: -1012; General average density:1.38 g/cm3;DIN regulations: Tubing (PVC-U) must meet DIN 6660/6661 regulations - DIN8061/62 Group B1 specification. Absorption of water during 24 hours:0.03%; Combustibility: Self-extinguishing; Difficult to inflame; along with necessary accessories Tube clips, tube clamps, special Glue, Special Cleaner, PVC sleeves Thread rod and nuts and bolts @ every 2 meters Etc. Factory Lot Test Report Have to submit by the Bidder
9 | Forwarding Tube (Transparent) 160 mm and Bends (Transparent): The forwarding tube should be made of uPVC of 160 mm OD x 153.6 mm ID. Good Physical tensile strength, general medium density, absorption of water, combustibility self-extinguishing. Colour Clear, Transparent.;DIN regulations: Tubing (PVC-U) must meet DIN 6660/6661 regulations - DIN8061/62 Group B1 specification.Absorption of water during 24 hours:0.03%; Combustibility: Self-extinguishing; Difficult to inflame . Factory Lot Test report Have to be submitted by the bidder
10 | Sleeves Where tubing passes through a wall, floor, ceiling or other barrier, the fire rating of the barrier is not accepted to be reduced. This must be achieved by equipping the system on these places, with crushing/self-extinguishing fire sleeves/collars that are certified by a notified body and are compliant with the local regulations. It should be made up of Polyvinylchloride PVC with length approx. 150 mm. The outer dia should be 168 mm & Inner dia should be 160 mm. Weight 0.43kg (approx.) factory quality and Test report for particular LOT have to be submit by the Bidder.
11 | Adhesive Glue: It can be used for bonding thermoplastic pressure piping systems made of rigid PVC. The boiling point should be 66°C (150.8 °F) and Explosion limit should be lower 1.36% (V) and upper 12.6% (V). It should be easy bonding, even of larger dimensions and clean handling. It should be excellent gap bridging capability.
12 | Diverter Units: The 3-way diverter units are switching device used at branching points in the system to direct the path of the carrier from a single tube at one end to three selectable tubes at the other end. The ends of the diverters are attached with forwarding tube network using steel clamps that are easily removable during servicing. To conform to EU Director for Electromagnetic tolerance 2004/108/EC; EEC directive for low voltage devices 93/68/EEC; EEC machines directors 98/37/EEC. Diverters are fitted with moveable tube that gently guides the carrier through the diverter in pre and auto-selected direction. The tube is positioned precisely at the selected port using proximity sensor built-in. The units are equipped with state-of-art maintenance free linear/circular gear drive mechanism & self-adjusting proximity digital sensors & noise-less operations.
13 | Carriers with RFID: With loading dimensions of 330mmx120mm. Special leak proof carriers with caps for secure use; fully integrated rubber seal; payload of 4 kg or more to carry Blood bags ,liquid ,blood samples etc.; Impact resistant and crystal clear polycarbonate middle body; With Loading Dimensions of 330mmH x 120mm ID; with Unique flip-top or Swivel caps for secure use; Metal-free design through rotation welding procedure; Impact resistant and crystal clear polycarbonate middle body, with integrated radio-frequency identification system( 2Each) . Each carrier should have soft sponge like foam inserts to hold samples securely.
14 | RFID The system should be provided with RFID as a standard solution, this helps in proper management of Carriers with the help of return of empty carrier. It also does not allow anything else to go in the system, but the carrier. All stations should have RFID reader Cards for the carrier, duly installed in the dispatch magazine. All the carriers should have RFID programmable chips on both sides of the carrier (2 per carrier)
15 | Control Data-Com-Power Composite Cable: Single composite control cable of high grade composite with grounding, power and data-com all three in one and fully screen and use grounded system concept to minimize electromagnetic and radio-frequency interference. Instrumentation wire type 2 x 2 x 0,22 x 1,9², Compatible with system on low voltage supply of 36V for human protection. All devices are connected with LVDC 36V using RS232 to RS485 platform using unicore cable based on RS485 connectivity which provides composite LV power supply and communication to all stations and devices in the network and hence are digital in nature and maintaining a real time live connectivity speed of minimum 70 – 120 devices per second which can be constantly monitored Live and Real Time on the PC monitor. Conforming to CE or EN standards along with necessary Tie-Strips for tying Control Cable Data -com-power with installed tubes & bends.
16 | Control Cabinet: Separate control cabinet for power supplies and all required control boards in the RPZ I (expandable up to 28 RP lines) according to valid standards and guidelines, incl. Cabinet lighting, forced ventilation, lockable door, incl. the necessary accessories.
17 | Power Pack: The entire system network work on 36-40 LVDC supply (except 3 phase blower) to enhance the network communication signal, speed and DC Voltage supplied at periodic lengths depending up on the network, 200 VA / 36 LVDC power & communication booster packs (power packs) are installed
18 | Degree of protection: IP54 same design as item 71.02.04 RPZ I: Control cabinet for electrical supply.
19 | Upgradability: The pneumatic tube system is designed and executed in such a way that changes as well extensions are possible at a later date and implemented systemically without additional efforts.
20 | Certification: Should be European CE or USFDA approved, where applicable
21 | Should have DIN certification, wherever applicable
22 | Should have certification to prove that the quoted and supplied tube meets International inflammable standards.
23 | Warranty: All the items except accessories like carriers etc. to be supplied with 2 years warranty.
24 | Prices of accessories/consumables like carriers, Velcro etc. will be covered for 5 years.